## Sheet1
Hazard ID | Regulation | Clause | Part/System Description | 42 Characteristic/ Intended use | Risk Analysis | [7 empty] | 5 Risk Evaluation |  | 6 Risk control
[5 empty] | 43 Hazard | 44 Estimate risk(s) for hazardous situation | [8 empty] | 62 Option Analysis |  | 63 Implementation of risk control measures | 64 Residual risk evaluation |  |  |  | 65 Benefit | 66 Other Hazards
[6 empty] | Normal/ Fault Condition (N/F)? | Initiating Events and Circumstances/Causes (Potential Cause of Hazard/ Failure) (원인)) | Foreseeable sequences of Events | Hazardous situation | Harm | Initial Risk Estimation
[11 empty] | Sev | Prob | Risk | Acceptable Risk?(Y/N) | Risk control Measurment (Description of Mitigation or Criteria) D: Inherent safety by design P: Protective measures in themedical device itself or in the manufacturing process I: Information for safety |  | Verification/ Validation | Sev | Prob | Residual Risk | Accptable Risk?(Y/N)
HZ-01 | IEC 62304 | 42 | Microprocessor | CameraInitalization | Function (Loss or deterioration of function) | N | Firmware upload failed | Not working equipment | Do not use equipment | Diagnostic error | 3 | 5 | 15 | N | D | Protective system | Software Validation report (6 Software Design Specification(SDS) ) | 2 | 2 | 4 | Y | N/A | No new and increased risks
P
I
HZ-02 | IEC 62304 | 42 | Microprocessor | ROM initalization | Function (Loss or deterioration of function) | N | Firmware upload failed | Not working equipment | Do not use equipment | Diagnostic error | 3 | 5 | 15 | N | D | Protective system | Software Validation report (6 Software Design Specification(SDS) ) | 2 | 2 | 4 | Y | N/A | No new and increased risks
P
I
HZ-03 | IEC 62304 | 42 | Memory device | Hard Disk | Function (Loss or deterioration of function) | N | Lack of capacity | Don't save information | Accurate information can not be obtained | Diagnostic error | 1 | 5 | 5 | Y | D | Warnings from the OS | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
I
HZ-04 | IEC 62304 | 42 | Memory device | ROM | Function (Loss or deterioration of function) | N | Lack of capacity | Not working equipment | Accurate information can not be obtained | Diagnostic error | 2 | 5 | 10 | N | D | Check file exists on the Hardware | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
I
HZ-05 | IEC 62304 | 42 | Sensor | Camera | Function (Loss or deterioration of function) | N | Don't connecte the Camera | inoperative the system | Accurate information can not be obtained | Diagnostic error | 2 | 5 | 10 | N | D | Designing software with reanalyze-function | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
I
HZ-06 | IEC 62304 | 42 | Sensor | Fiber | Function (Loss or deterioration of function) | N | Don't connecte the Fiber | inoperative the system | Accurate information can not be obtained | Diagnostic error | 1 | 5 | 5 | Y | D | Designing software with reanalyze-function | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
I
HZ-07 | IEC 62304 | 42 | Energy source | Camera | Function (Loss or deterioration of function) | N | Don't connecte the Camera | inoperative the system | Accurate information can not be obtained | Diagnostic error | 2 | 5 | 10 | N | D | Designing software with reanalyze-function | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
I
HZ-08 | IEC 62304 | 42 | Safety features | OverCurrent | Function (Loss or deterioration of function) | N | Too high/low Current Value | System: instability | Accurate information can not be obtained | Diagnostic error | 3 | 5 | 15 | N | D | Production overcurrent protection circuit | Software Validation report (6 Software Design Specification(SDS) ) | 2 | 2 | 4 | Y | N/A | No new and increased risks
P
I
HZ-09 | IEC 62304 | 42 | Safety features | Noise | Function (Loss or deterioration of function) | N | Too high Noise Value | Noise displays | Accurate information can not be obtained | Diagnostic error | 3 | 5 | 15 | N | D | Noise removal filter configuration | Software Validation report (6 Software Design Specification(SDS) ) | 2 | 2 | 4 | Y | N/A | No new and increased risks
P
I
HZ-10 | IEC 62304 | 42 | Communications | Save : Image | Function (Loss or deterioration of function) | N | User is not checking Memory | Failed to save the image | Save failed | Diagnostic error | 1 | 5 | 5 | Y | D | The protective earth terminals implementation Performin a safety test | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
I
HZ-11 | IEC 62304 | 42 | Communications | Save : Cine | Function (Loss or deterioration of function) | N | User is not checking Memory | Failed to cine the image | Save failed | Diagnostic error | 1 | 5 | 5 | Y | D | An error message is displayed | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
I
HZ-12 | IEC 62304 | 42 | Communications | Rotation | Function (Loss or deterioration of function) | N | User is not Rotation | Failed to rotation the image | Do not run (Rotation) | Diagnostic error | 1 | 5 | 5 | Y | D | An error message is displayed | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
I
HZ-13 | IEC 62304 | 42 | Communications | Brightness | Function (Loss or deterioration of function) | N | User is not Brightness | Failed to brigtness the image | Do not run (Brigtness) | Diagnostic error | 1 | 5 | 5 | Y | D | An error message is displayed | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
I
HZ-14 | IEC 62304 | 42 | Communications | Folder | Function (Loss or deterioration of function) | N | User is not checking Folder | Find the folder fails | Do not run (Folder) | Diagnostic error | 1 | 5 | 5 | Y | D | An error message is displayed | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
I
HZ-15 | IEC 62304 | 42 | Communications | Zoom | Function (Loss or deterioration of function) | N | User is not Zoom | Don't operate the Zoom | Do not run (Zoom) | Diagnostic error | 1 | 5 | 5 | Y | D | Do not expand anymore, and collapse | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
I
HZ-16 | IEC 62304 | 42 | Imaging Processing and Motion (H/W) | ADC | Function (Loss or deterioration of function) | N | Firmware: Aliasing occurs | System: instability | Accurate information can not be obtained | Diagnostic error | 2 | 5 | 10 | N | D | function to check the verification cacluation | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
I
HZ-17 | IEC 62304 | 42 | Interface Requirements | PC / Monitor | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the Monitor | Do not use the PC / Monitor | Diagnostic error | 2 | 5 | 10 | N | D | Well put power protection terminal Connector checks | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
I
HZ-18 | IEC 62304 | 42 | Interface Requirements | Mouse | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the Mouse | Do not use the Mouse | Diagnostic error | 2 | 5 | 10 | N | D | Well put power protection terminal Connector checks | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
| IEC 62304 |  | Interface Requirements | Mouse |  | N | Miss the connection | Don't operate the Mouse | Do not use the mouse |  |  |  |  | Y | I
HZ-19 | IEC 62304 | 42 | Interface Requirements | Keyboard | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the Keyboard | Do not use the keyboard | Diagnostic error | 1 | 5 | 5 | Y | D | Well put power protection terminal Connector checks | Software Validation report (6 Software Design Specification(SDS) ) | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
[7 empty] | Miss the connection |  | Do not use the mouse | [5 empty] | I
HZ-20 | IEC 62304 | 42 | Interface Requirements | USB | Function (Loss or deterioration of function) | N | USB Reading Time | Don't operate the USB | Do not use the USB | Diagnostic error | 3 | 5 | 15 | N | D | Well put power protection terminal Connector checks | Software Validation report (6 Software Design Specification(SDS) ) | 2 | 2 | 4 | Y | N/A | No new and increased risks
P
I
HZ-21 | IEC 62304 | 42 | Timing and memory requirements | System Boot Time | Function (Loss or deterioration of function) | N | Memory initial failed | Delayed loading speed | Accurate information can not be obtained | Diagnostic error | 2 | 5 | 10 | N | D |  |  | 1 | 2 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Program standby status | Software Validation report (6 Software Design Specification(SDS) )
I
HZ-22 | IEC 62304 | 42 | Identification of off-the-Shelf Software | Version | Function (Loss or deterioration of function) | N | Can not display version | Not displayed | Not Shown (Version) | Diagnostic error | 1 | 5 | 5 | Y | D |  |  | 1 | 2 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (6 Software Design Specification(SDS) )
I
HZ-23 | IEC 62304 | 42 | Imaging Processing and Motion (S/W) | Brightness | Function (Loss or deterioration of function) | N | RF data input: brightness values fail | Wrong output the screen | Brightness should not be adjusted | Diagnostic error | 1 | 5 | 5 | Y | D |  |  | 1 | 2 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (6 Software Design Specification(SDS) )
I
HZ-24 | IEC 62304 | 42 | Imaging Processing and Motion (S/W) | Color Reconstruction (RGGB) | Function (Loss or deterioration of function) | N | RF data input: RGGB values fail | Wrong calculate the RGGB | RGGB should not be adjusted | Diagnostic error | 2 | 5 | 10 | N | D |  |  | 1 | 2 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (6 Software Design Specification(SDS) )
I
HZ-25 | IEC 62304 | 42 | Imaging Processing and Motion (S/W) | Rotation | Function (Loss or deterioration of function) | N | RF data input: Rotation values fail | Wrong calculate the Rotation flag | Rotation should not be adjusted | Diagnostic error | 2 | 5 | 10 | N | D |  |  | 1 | 3 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (6 Software Design Specification(SDS) )
I
HZ-26 | IEC 62304 | 42 | Imaging Processing and Motion (S/W) | AutomaticExposureControl | Function (Loss or deterioration of function) | N | RF data input: AEC values fail | Do not want to AEC | AEC should not be adjusted | Diagnostic error | 2 | 5 | 10 | N | D |  |  | 1 | 4 | 4 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (6 Software Design Specification(SDS) )
I